## Slide 1: Fluoroscopy Operation and Safety: Influence on Patient Dose

## Slide 2: Fluoroscopy
High dose procedures
Last resort after routine radiography
100-500 mA for routine studies vs. an average of 3mA for fluoroscopy*
Dynamic studies
Transient structures
Fluorescence

## Slide 4: Standard X-Ray Tube
Rotating anode
Small focal spot
Operates in the range of 0.5 to 5 mA

## Slide 5: Milliamperage (mA)
mA controls density and photon quantity
Directly proportional to dose
Typical mA’s:
  Average of 3 mA
  Spot imaging > 100 mA

## Slide 6: Kilovoltage Peak (kVp)
Contrast
Differential absorption
Quality
Peak kVp

## Slide 7: Collimation
Principle method to reduce patient dose
Image will NOT be brighter if collimation is open wider

## Slide 8: Collimation and Title 17
With automatic collimating devices on a fluoroscopic unit, an unexposed border must be visible at all heights above the table
With manual collimation: 14 inches above table

## Slide 9: Filtration
Hardens the beam
Inherent = 0.5mm AL
Added = 2.0mm AL
Total minimum filtration is 2.5 mm/AL

## Slide 10: Source to Table Top Distance (STTD)
Should be 18”
Shall be no less than 12”
Also called target-to-panel distance (TPD)

## Slide 11: Object to Image Distance
OID in fluoroscopy refers to the distance between the patient and the image intensifier
Minimize at all times
Ultimately reduces patient dose (to be discussed later)

## Slide 13: Low Absorption Table Tops
Made of composite materials like Bakelite, fiberglass, polycarbonate, carbon fiber, etc.
Table cannot excessively absorb the primary beam
Shall be no more than 1 mm Al (aluminum equivalent) at 100 kVp

## Slide 14: Exposure Switch
Must be of the dead-man type
Exposure is discontinued when pressure is released

## Slide 15: Primary Protective Barrier
Refers to the image intensifier tube housing
2 mm Pb at 125 kVp
Primary barrier for room
  1/16” Pb at 6’8”
Secondary barrier for room
  1/32” Pb at 6’8”

## Slide 16: Bucky Slot Cover
0.25 mm Pb or equivalent
For operator protection

## Slide 17: Protective Drapes
0.25 mm Pb or equivalent
Prevents scatter from reaching the operator
Theoretically, one could receive 500 mR / hr at 1 foot from the table
The technologist/operator only gets 1/1000 of what the patient receives

## Slide 18: Exposure Time
Fluoroscopy exposure is pulsed rather than continuous
System must have last image hold (LIH)
Want to limit the “beam on” time

## Slide 19: Exposure Rates
Air Kerma Rate (AKR)
For normal operation, up to 5 R / min OR 2.2 R/mA to Max 5 mA
10 R / min with a coupled recording device
44 mGy/min = 5 R/min
88 mGy/min = 10 R/min

## Slide 20: Reset Timer
Preset to 5 minutes
Does NOT terminate exposure
Alerts fluoroscopist to accumulated fluoroscopy time
If an alarm sounds, it must be for a minimum of 2 seconds

## Slide 21: Illumination (Room)
Room must have low illumination
Enhances black and white vision
Indirectly impacts patient dose

## Slide 22: Quantum Mottle
Too few incident photons at the input phosphor
Appears as noise or “snow”
Poor image quality
More evident using high contrast or magnification modes

## Slide 23: Resolution
Pertains to detail of small objects adjacent to each other
Line pair tester
Line pair / mm (lp/mm)
Modulation Transfer Function (MTF)
Ideal value = 1
In reality, MTF < 1

## Slide 24: Contrast
Subject contrast
Controlled by kVp
Short and long scale
Detector contrast ratio 15:1
Decreases with age

## Slide 25: Contrast
The brightness ratio of the periphery of the output screen to the center of the output screen is referred to as contrast

## Slide 26: Distortion
Size distortion
OID
Shape distortion
Geometric differences between input and output phosphors
Pincushion effect

## Slide 27: Vignetting
Also due to geometric differences between input and output phosphors
Fall-off of brightness at the periphery of the image

## Slide 28: Cameras & Lag
Image build-up and decay in the TV camera
Vidicon camera:
  high lag, low dose
Plumbicon camera:
  Low lag, high dose
Cameras are attached at the output phosphor
NOTE: TV cameras are NOT used with CCD or flat panel detectors

## Slide 29: Synchronization of Coupled Device
A coupled device is a camera or similar
Operates at the same frequency of x-ray pulses
Shutters closed when film is transported

## Slide 31: Framing (Pulse) Frequency
Always in division of 60
Typically 30 frames/pulses per second
Higher framing frequency means higher dose to the patient

## Slide 32
Exact Framing
Coupled Device
Output Phosphor

## Slide 33
Under Framing

## Slide 34
Over Framing

## Slide 35
Total Over Framing

## Slide 36: F Number
References the speed of a given camera
The lower the number, the faster the camera
Low number will reduce patient dose
F= Focal length / diameter of lens

## Slide 37: Recording Devices
Conventional spot film
  Highest dose & best quality
Spot film camera (photospot, digital photospot)
  Low dose with lower quality
Video Tape
  Lowest dose with poor quality
Digital recording
  Lower dose and good quality

## Slide 38: Gonadal Shielding
Mandated by law
California requires a minimum of 0.5 mm Pb
“Shall be no less than 0.5 mm Pb”

## Slide 39: Grids
Use always increases patient dose
Most fluoroscopy systems have low ratio grids
Avoid using grids in pediatric patients

## Slide 40: Body Habitus
Tissue density
Additive conditions
Destructive conditions
High atomic number and attenuation

## Slide 41: Scatter and Absorption
High kVp
Large field size
Thick body habitus
Extended source of radiation

## Slide 42: X-Ray Generators
Common generators include single phase, three-phase, medium, and high frequency
Three-phase and high frequency have certain advantages such as constant potential and high mA resulting in unlimited exposure times

## Slide 43: Contrast Media

## Slide 44: Contrast Media
The purpose of contrast media is to alter the characteristics of the patient’s body
Many structures will not be apparent unless contrast media has been introduced
Contrast media will either be positive or negative

## Slide 45: Contrast Media (cont)
Negative contrast media are those which are considered radiolucent
  These media would consist of room air, CO2, or something similar
Positive media absorb x-rays and are considered radiopaque
  Barium, iodine, or something with a higher atomic number
Must be non-toxic

## Slide 46: BaSO4
Most frequently chosen contrast
White, crystalline powder that is mixed with water to make a suspension
Can be administered by mouth, rectum, and tube insertion
Relatively nontoxic, however, if there is a leakage into the peritoneal cavity or blood stream, an adverse reaction can occur
Can be constipating

## Slide 47: Barium Preparation

## Slide 50: Iodinated Contrast
Iodine has an atomic number of 53
  Absorbs photons through photoelectric interactions
Most are water-based (in solution)
Can be ingested (for GI studies) and injected (venous, arterial, joint spaces, thecal sac, etc.)

## Slide 52
Retrograde Pyelogram

## Slide 55: Ionic Media
All iodinated contrast media are composed of a cation (+) and an anion (-)
The cation is either the sodium or meglumine compound in ionic media
The ions dissociate (ionize) completely once inside the body

## Slide 56: Ionic Media
Most are higher-osmolality contrast media (HOCM)
Osmolality is a measure of the total number of particles in solution per kilogram of water
Osmotic pressure controls fluid movement in the body

## Slide 57: Ionic Media (cont)
HOCM changes the blood plasma
Pulls water from the tissue into the blood stream causing an electrolyte imbalance
Movement also causes hypervolemia and blood vessel dilatation

## Slide 58: Electrolytes
Anion
  Negatively charged constituent
Cation
  Positively charged constituent
Fe+
Cl-
H2O
Na+
NaCl
K+
Albumin
H2O
Cl-
Fe+
Fe+
H2O
H+
H+
OH-
OH-

## Slide 59: Effects of Hypertonic Contrast Media
Water moves into lumen to dilute environment
Net effect an increase in the volume of fluid in the lumen, but homeostasis is restored
Fe+
Cl-
H2O
Na+
NaCl
K+
Albumin
H2O
H2O
Cl-
Fe+
Fe+
H2O
H2O
H2O
OH-
H+

## Slide 60
Effects of Hypotonic Contrast Media
Fe+
Cl-
H2O
Na+
NaCl
K+
Albumin
H2O
H2O
Cl-
Fe+
Fe+
H2O
H2O
H2O
Water moves out to equalize concentration
Net reduction in intracellular volume
Edema in surrounding tissue

## Slide 61: Non-ionic Media
Contrast media was developed over time to reduce side effects
Some media do not dissociate into cations and anions
Others are still ionic but with large molecules which do not cause osmotic effects

## Slide 62: Non-ionic Media (cont)
Lower-osmolality contrast media (LOCM)
Cations were removed without any loss of diagnostic information
Replaced with compounds that do not dissociate
Osmolality closer to human plasma
LOCM are more water soluble, may be less likely to trigger allergic effects

## Slide 63: Effects with Non-Ionic Contrast Media
Less net movement of ions in either direction
Less change in electrolyte balance
Minimal fluid shift through kidneys
Fe+
Cl-
H2O
Na+
NaCl
K+
Albumin
H2O
H2O
Cl-
Fe+
Fe+
H2O
H2O
H2O

## Slide 64: All Iodinated Contrasts
Iodine in its natural state is chemically reactive and can be toxic in the body
Both types of iodinated media contain additives (citrate, calcium disodium edetate)
Prevent iodine atoms from being removed from the contrast molecules

## Slide 65: Patient Assessment
Patient history must be taken for information pertinent to contrast injection and anything that may predict an allergic reaction
Prevention of serious adverse effects
Correct medical treatment must be initiated immediately in the event of a reaction

## Slide 66: Effects of Iodinated Contrast
General effects
  HOCM show greater effects and adverse reactions
  Viscosity is influenced by concentration and size of the molecule
  Affects ability to inject

## Slide 67: Renal Effects
HOCM cause arteries of the kidneys to expand because of the osmotic effect
Expansion results in release of vasoconstrictors to constrict the renal arteries
End result is diminished blood supply to the kidneys

## Slide 68: Renal Effects (cont)
Osmotic effects presumed to increase the amount of molecular substances reabsorbed by the renal tubules
Osmotic diuresis and dehydration
BUN and creatinine are good indicators for possible contrast media-induced renal effects

## Slide 69
http://www.ivy-rose.co.uk/Topics/Urinary_System_Nephron_Diagram.htm

## Slide 70: Patient Reactions
Expected side effects
Mild (minor)
Moderate
Severe
Vasovagal

## Slide 71: Expected Side Effects
A feeling of flushing or warmth
Nausea and / or vomiting
Headache
Pain at the injection site
  Check for extravasation
Altered taste, may be metallic

## Slide 72: Mild Adverse Reaction
Nausea / vomiting
Cough
Feeling of warmth
Headache
Dizziness
Shaking
Itching
Pallor

## Slide 73: Moderate Adverse Reaction
Tachy- or bradycardia
Hyper- or hypotension
Dyspnea
Bronchspasm or wheezing
Patient complaint of feeling of throat closing (laryngeal edema)

## Slide 74: Severe Anaphylactic Reaction
Dyspnea related to laryngeal edema
Hypotension
Seizures
Cardiac arrythmia
Lack of patient response
Cardiac / respiratory arrest

## Slide 75: Vasovagal Response
Reaction to the procedure itself than the contrast
Due to high anxiety on the part of the patient
Pallor (paleness), cold sweats, rapid pulse, syncope (or complaining of feeling faint), bradycardia, hypotension
Stop contrast, place patient flat or in Trendelenberg position, notify supervising physician, stay with the patient

## Slide 76: Your Role
Responsibility of patient comfort throughout the duration of the procedure
Identify signs and symptoms of adverse reactions
Obtain a thorough patient history that can indicate contrast media contraindications or increased possibility of adverse reactions