- Technical Aspects to the Fetal Exam

- Michelle Wilson  Ed.D(c), RDMS, RDCS

- ALARA

- “As low as reasonably achievable”

- Affected by the duration and level of exposure

- Measured in sonography by Thermal Index and Mechanical Index

- Bio Effects

- Mechanical energy produces a pressure wave

- Reflection back to transducer forms an image

- Physical effects:

- Thermal

- Cavitational

- Thermal effects

- Ultrasound wave absorbed

- Energy converted to heat

- High in bone, low in fluids

- Heat determined on:

- Tissue characteristics

- Perfusion, conduction of heat

- Ultrasound intensity

- Length of time scanned

- Intensity

- Dependent on power output and position of tissue in the beam profile

- Intensity can be altered

- Power

- Mode

- Scan depth

- Focus

- Zoom

- Sample volume

- Cavitation

- Formation of transient or stable bubbles

- Inertial cavitation

- Gas-filled cavity grows and contracts

- Collapse of the bubble can cause increased temperatures and pressures

- Effects on Fetuses

- Rise in temperature

- Control exposure time, power output

- Use controls that don’t increase power and intensity

- Points to Ponder

Control exposure to patient

Good image-reduced output power

Gain, Post processing, dynamic range

Minimize the use of spectral Doppler

- Sonographic Techniques

- High standards

- Highest  Frequency

- Perpendicular incidences when possible

- Shortest route

- Best acoustic window

- High Frame Rate

- Focal Zones

- Basic Doppler Principles

- Measurements of movement

- Pulses transmitted to detect movement

- Echoes from moving scatterers exhibit differences

- Measured as phase shift from Doppler frequency

- Doppler Shift

- Motion in direction of beam

- Signal dependent on:

- Velocity

- Frequency

- Angle

- Continuous Wave and Pulsed Wave

- Continuous transmission of sound beam

- Location ambiguity encountered

- Not utilized in fetal echo

- Pulsed wave

- Allows measurement of depth

- Sample volume manipulation

- Pulsed Wave Doppler

- Limited zone of sampling along the beamline

- Can and will cause heating

- Use sparingly

- Factors Affecting Pulsed Wave

- Direction of beam

- Sample volume

- Angle correction

- Filter

- Pulsed Doppler

- Optimizing

- Power

- Transmitted power into tissue

- Gain

- Overall sensitivity to flow signals

- PRF

- Scale

- Gate size

- Smaller the gate, more sensitive

- Beam steering

- Improved flow angle more accurate velocity calculation

- Color Doppler

- Fewer shorter pulses

- Mean frequency shift-displayed as a color pixel

- Variance

- Transducer elements share between B-mode and color flow imaging

- Pulses for color are 3-4 times longer than B mode

- Decrease axial resolution

- Color Flow Imaging

- Identify vessels

- Presence and direction of flow

- Circulation anomalies

- FACTORS AFFECTING THE COLOR FLOW IMAGE

- Frequency

- Higher-more sensitivity to low flow, better spatial resolution

- Lower-better penetration

- Velocity scale

- Low PRF used for low velocitites

- Region of interest

- Reduce width and maximum depth

- Improve frame rate improved spatial resolution

- Focus

- Level of interest

- Power Doppler

- Energy, Amplitude, Angiography

- Magnitude of the color flow output displayed

- Does not display flow direction or velocities

- Sensitive to extra noise

- Motion-Mode (M-Mode)

- Analyzing moving structures

- Function of:

- Depth

- Time

- High sampling ability

- Summary

- Practice ALARA

- Be aware of technical parameters we can change

- Limit exposure

- Especially with EV studies

- Especially with 1st trimester studies